## Slide 1
Automated, dynamic, SI-traceable field calibration of atmospheric elemental and oxidized mercury
Authors: Seth Lyman, Colleen Jones, KarLee Zager, Trevor O’Neil

## Slide 2
Acknowledgements
Funding provided by the
U.S. National Science Foundation
Patents related to this work
10012622
10545123
Docket P23018.02 (pending)

## Slide 3
Hg0 calibration:
The state of the science

## Slide 4
Hg0 calibration:
NIST and others have primary standards, but no ambient-level field-ready standards
The state of the science

## Slide 5
Hg0 calibration:
NIST and others have primary standards, but no ambient-level field-ready standards
Field instruments are calibrated with syringe injections from a saturated vapor source
The state of the science

## Slide 6
Hg0 calibration:
NIST and others have primary standards, but no ambient-level field-ready standards
Field instruments are calibrated with syringe injections from a saturated vapor source
  Accuracy and stability challenges
  Not SI-traceable
  Few calibrations performed in practice
The state of the science

## Slide 7
HgII calibration:
The state of the science

## Slide 8
HgII calibration:
No commercial options for ambient levels
The state of the science

## Slide 9
HgII calibration:
No commercial options for ambient levels
Lack of calibration led to ~20 years of resources wasted on biased KCl denuder-based HgII measurements
The state of the science

## Slide 10
We are putting our effort
into permeation tube methods
First used by Landis et al., 2002?

## Slide 11
Capable of low, continuous emission rates
Shown to work in real field conditions
Simple (Tekran 2537 uses permeation tubes for internal calibrations, but it is neither independent nor SI-traceable)
We are putting our effort
into permeation tube methods

## Slide 12
Basic principles of our method

## Slide 13
SI-traceable gravimetric verification of permeation rate
Elgiar et al., 2024

## Slide 14
Independent determination of SI traceability
(Elgiar et al., 2024)

## Slide 15
Independent determination of SI traceability
(Elgiar et al., 2024)

## Slide 16
Two remaining problems

## Slide 17
Two remaining problems
HgII permeation rates are too high
  Lowest we can do is ~100X ambient levels

## Slide 18
Two remaining problems
HgII permeation rates are too high
  Lowest we can do is ~100X ambient levels

## Slide 19
Two remaining problems
HgII moves slowly through all tubing materials
  1-2 hr tubing equilibration times needed

## Slide 20
Two remaining problems
What our automated system tried
(and failed at):

## Slide 21
Two remaining problems
What our automated system tried
(and failed at):
10-sec injection

## Slide 22
Two remaining problems
What our automated system tried
(and failed at):
10-sec injection
140-sec pause

## Slide 23
Two remaining problems
What our automated system tried
(and failed at):
10-sec injection
140-sec pause
10-sec injection

## Slide 24
Two remaining problems
What our automated system tried
(and failed at):
10-sec injection
140-sec pause
10-sec injection
140-sec pause

## Slide 25
Two remaining problems
What our automated system tried
(and failed at):
10-sec injection
140-sec pause
10-sec injection
140-sec pause
We overcame that problem by manually inserting into inlet

## Slide 26
Two remaining problems
Step 1: Equilibrate for 2 hours by flowing Hg output to ambient air

## Slide 27
Two remaining problems
Step 2: Manually insert calibrator outlet into analyzer for 10 sec

## Slide 28
Two remaining problems
Step 3: Vent for 140 sec

## Slide 29
Two remaining problems
Repeat steps 2 and 3 over and over

## Slide 30
Two remaining problems
But this is neither routine nor automated, which was our goal

## Slide 31
New system: dynamic draw-off

## Slide 32
New system: dynamic draw-off

## Slide 33
New system: dynamic draw-off
MFC to deliver inert gas

## Slide 34
New system: dynamic draw-off
MFC to deliver inert gas
Another MFC to draw off a portion of the Hg+inert gas

## Slide 35
New system: dynamic draw-off
MFC to deliver inert gas
Another MFC to draw off a portion of the Hg+inert gas
Custom, ultra-inert venturi flow measurement to measure remaining flow

## Slide 36
New system: dynamic draw-off
MFC to deliver inert gas
Another MFC to draw off a portion of the Hg+inert gas
Custom, ultra-inert venturi flow measurement to measure remaining flow
Remaining flow
Total original
flow
Fractional
dilution rate
=

## Slide 37
New system: dynamic draw-off

## Slide 38
The system appears to be working

## Slide 39
~10-min equilibration time
The system appears to be working

## Slide 40
~10-min equilibration time
Mostly HgII
The system appears to be working

## Slide 41
The system appears to be stable
at different permeation rates
1.9% of original flow remaining
0.25 pg/s effective calibration rate
3.0% of original flow remaining
0.40 pg/s effective calibration rate

## Slide 42
Remaining challenge with the new system

## Slide 43
Remaining challenge with the new system
Small drifts in flow measurements lead to unacceptable uncertainty in output.
Lots of flow measurements are underway to resolve this.

## Slide 44
Where we are now
Hg0

## Slide 45
Where we are now
Hg0
SI-traceable

## Slide 46
Where we are now
Hg0
SI-traceable
Field ready

## Slide 47
Where we are now
Hg0
SI-traceable
Field ready
Automated

## Slide 48
Where we are now
Hg0
SI-traceable
Field ready
Automated
Dynamically adjusted calibration amount

## Slide 49
Where we are now
HgII

## Slide 50
Where we are now
HgII
SI-traceable

## Slide 51
Where we are now
HgII
SI-traceable
Field ready

## Slide 52
Where we are now
HgII
SI-traceable
Field ready
Automated

## Slide 53
Where we are now
HgII
SI-traceable
Field ready
Automated
Dynamically adjusted calibration amount

## Slide 54
Let’s talk
Seth Lyman, PhD
Utah State University
seth.lyman@usu.edu
(435) 630-1433